## OD PMT replacement 2005-6
OD PMT Replacement List for SK Reconstruction 2005/6
|  | 2006-02-20 18:51 | HGB
# | SL | PMT | cable# | paddle# | HV# | HV(V) | zener | dead since | replaced | s/n | HV_PMT | HV_diff | tested | uA_test | R(M) | Type | failure notes
1 | 18 | T1P08P01A | 21441 | 1.2.08.09 | 1.2.03 | 1820 |  | 2003-04-01 | 2005-11-02 | SA4111 | 1820 | 0 | 2005-11-03 |  | - | IMB | high current, no signal
2 | 18 | T2M01P02A | 21465 | 2.2.08.11 | 2.2.03 | 1620 |  | 2005-10-15 | 2005-11-02 | SA4142 | 1689 | -69 | 2005-11-03 | 64 | 24.1125 | Rec | hvtrip, then 600 uAmp
3 | 18 | T2M01P03B | 21510 | 2.2.10.04 | 2.2.05 | 1610 |  | 2003-03-05 | 2005-11-02 | SA4205 | 1687 | -77 | 2005-11-03 | 61.4 | 25.0215 | Upg | hvtrip
4 | 18 | T2M04P01A | 21430 | 2.2.08.01 | 2.2.03 | 1620 |  | 2005-02-27 | 2005-11-02 | SA4102 | 1613 | 7 | 2005-11-03 | 62.1 | 24.887 | IMB | flasher; killed QTC daughtercard, too
5 | 18 | T3M01M01B | 21417 | 3.2.07.09 | 3.2.02 | 1710 |  | 2004-08-31 | 2005-11-02 | SA4141 | 1709 | 1 | 2005-11-03 | 61.4 | 26.6502 | IMB | flasher candidate
6 | 18 | T3M01M02A | 21370 | 3.2.08.12 | 3.2.03 | 1760 |  | 2004-11-25 | 2005-11-02 | SA4132 | 1757 | 3 | 2005-11-03 | 62.9 | 26.7809 | IMB | hvtrip
7 | 18 | T3M03M02A | 21368 | 3.2.09.04 | 3.2.04 | 1690 |  | 2005-12-08 |  | SA4110 | 1843 | -153 |  |  | - | Rec | hvtrip (new!)
8 | 18 | T3M04M03A | 21337 | 3.2.10.10 | 3.2.05 | 1580 | 200 | 2005-11-01 | 11/30 (?) | SA4206 | 1700 | -120 | 12/8 (?) | 59 | 25.5797 | IMB | noisy, low gain
9 | 18 | T3M06M03B | 21350 | 3.2.11.03 | 3.2.06 | 1600 |  | 2005-05-02 | 2005-11-02 | SA4127 | 1576 | 24 | 2005-11-09 | 63 | 24.1968 | IMB | hvtrip
10 | 18 | T3M06M04A | 21308 | 3.2.12.05 | 3.2.07 | 1640 |  | 2003-09-17 | 2005-11-02 | SA4131 | 1649 | -9 | 2005-11-09 | 59 | 26.5966 | IMB | hvtrip
11 | 18 | T4P03M01A | 21404 | 4.2.08.01 | 4.2.03 | 1640 |  | 2005-07-25 | 2005-11-02 | SA4207 | 1648 | -8 | 2005-11-09 | 58.5 | 26.8342 | IMB | flasher candidate
12 | 18 | T4P03M03B | 21358 | 4.2.10.10 | 4.2.05 | 1640 |  | 2002-12-23 | 2005-11-02 | SA4108 | 1695 | -55 | 2005-11-09 | 58.5 | 26.8342 | Rec | hvtrip
13 | 18 | T4P08M02B | 21393 | 4.2.10.04 | 4.2.05 | 1640 |  | 2005-04-20 | 2005-11-02 | SA4203 | 1699 | -59 | 2005-12-08 | 58.5 | 26.8342 | IMB | flasher; killed QTC daughtercard, too
14 | 17 | W2P08P15B | 20510 | 2.1.16.12 | 2.1.03 | 1740 |  | 2005-10-26 | 2005-11-03 | SA4104 | 1764 | -24 | 2005-11-03 | 62.1 | 26.8193 | Upg | hvtrip
15 | 17 | W2P08P18B | 20612 | 2.2.05.06 | 2.2.00 | 1795 |  | 2005-12-08 | 2005-12-14 | SA4163 | 1797 | -2 | 2005-12-16 | 64.2 | 26.7595 | IMB | hvtrip (new!)
16 | 17 | W3P08P25B | 20850 | 3.1.18.10 | 3.1.05 | 1790 | 200 | 2004-08-27 | 2005-11-03 | SA4154 | 1792 | -2 | 2005-11-03 | 64 | 26.7688 | IMB | hvtrip
17 | 17 | W3P08P28A | 20935 | 3.2.07.03 | 3.2.02 | 1710 |  | 2003-01-30 | 2005-11-03 | SA4133 | 1711 | -1 | 2005-11-03 | 61.4 | 26.6502 | Upg | hvtrip
18 | 16 | W1P07P06B | 20203 | 1.1.17.10 | 1.1.04 | 1800 |  | 2003-12-23 | 2005-12-03 | SA4211 | 1812 | -12 | 2005-12-16 | 64 | 26.925 | IMB | hvtrip, arcing (QTC killed)
19 | 16 | W1P07P08A | 20254 | 1.2.03.05 | 1.1.10 | 1700 | 150 | 2005-10-28 | 2005-12-03 | SA4137 | 1707 | -7 | 2005-12-16 | 61 | 26.6689 | IMB | hvtrip
20 | 16 | W3P07P19B | 20645 | 3.1.01.08 | 3.0.00 | 1615 |  | 2005-11-01 | 2005-12-03 | SA4126 | 1707 | -92 | 2005-12-16 | 59.3 | 26.0344 | IMB | hvtrip (new!)
21 | 16 | W3P07P22B | 20747 | 3.1.10.02 | 3.0.09 | 2060 | 200 | 2003-09-30 | 2005-12-03 | SA4109 | 2067 | -7 | 2005-12-16 | 73.8 | 26.7133 | IMB | hvtrip
22 | 16 | W4P07P35B | 21172 | 4.1.18.12 | 4.1.05 | 1980 |  | 2002-10-16 | 2005-12-03 | SA4251 | 1985 | -5 | 2005-12-16 | 71 | 26.6873 | Rec | hvtrip, bad cable?
23 | 15 | W2P06P15B | 20508 | 2.1.16.08 | 2.1.03 | 1740 |  | 2005-12-08 | 2005-12-09 | SA4103 | 1747 | -7 | 2005-12-16 | 62.1 | 26.8193 | IMB | hvtrip (new!)
24 | 15 | W3P06P26A | 20865 | 3.2.01.03 | 3.1.08 | 1980 |  | 2004-06-02 | 2005-12-09 | SA4179 | 1984 | -4 | 2005-12-16 | 70.6 | 26.8453 | IMB | hvtrip (resistance ok at low voltages)
25 | 15 | W4P06P36B | 21205 | 4.2.01.08 | 4.1.08 | 1640 |  | 2005-08-15 | 2005-12-09 | SA4120 | 1727 | -87 | 2005-12-16 | 58.5 | 26.8342 | IMB | hvtrip
26 | 14 | W1P05P01A | 20014 | 1.1.03.03 | 1.0.02 | 1690 |  | 2006-02-17 |  | SA4121 | 1800 | -110 |  |  |  | IMB | hvtrip
27 | 14 | W1P05P03B | 20099 | 1.1.08.12 | 1.0.07 | 1995 |  | 2003-06-09 | 2005-12-14 | SA4161 | 1998 | -3 | 2005-12-16 | 71.3 | 26.7804 | Upg | hvtrip
28 | 14 | W1P05P05A | 20150 | 1.1.14.07 | 1.1.01 | 1775 |  | 2002-10-09 | 2005-12-14 | SA4149 | 1775 | 0 | 2005-12-16 | 62.9 | 27.0194 | IMB | hvtrip
29 | 14 | W3P05P23B | 20779 | 3.1.12.08 | 3.0.11 | 1920 |  | 2003-03-16 | 2005-12-14 | SA4237 | 1919 | 1 | 2005-12-16 | 68.5 | 26.8292 | IMB | hvtrip
30 | 14 | W4P05P30B | 21000 | 4.1.04.06 | 4.0.03 | 1880 |  | 2002-11-22 | 2005-12-14 | SA4159 | 1882 | -2 | 2005-12-16 | 67.3 | 26.7346 | IMB | hvtrip
31 | 14 | W4P05P33B | 21102 | 4.1.12.12 | 4.0.11 | 1860 |  | 2004-11-17 | 2005-12-14 | SA4115 | 1860 | 0 | 2005-12-16 | 66.2 | 26.8967 | Rec | hvtrip
32 | 13 | W1P04P02B | 20064 | 1.1.05.12 | 1.0.04 | 1860 | 100 | 2005-01-21 | 2005-12-19 | SA4143 | 1866 | -6 | 2006-02-13 | 66.1 | 26.9392 | IMB | hvtrip
33 | 13 | W1P04P03A | 20081 | 1.1.08.09 | 1.0.07 | 1995 |  | 2004-09-15 | 2005-12-19 | SA4195 | 1998 | -3 | 2006-02-13 | 71.3 | 26.7804 | Upg | hvtrip
34 | 13 | W2P04P14B | 20472 | 2.1.13.06 | 2.1.00 | 1750 |  | 2002-12-09 | 2005-12-19 | SA4119 | 1754 | -4 | 2006-02-13 | 62 | 27.0258 | IMB | hvtrip
35 | 13 | W3P04P21A | 20693 | 3.1.06.09 | 3.0.05 | 2030 |  | 2004-02-15 | 2005-12-19 | SA4255 | 2030 | 0 | 2006-02-13 | 72.7 | 26.723 | IMB | hvtrip
36 | 13 | W3P04P23B | 20778 | 3.1.12.06 | 3.0.11 | 1920 |  | 2005-08-26 | 2005-12-19 | SA4180 | 1930 | -10 | 2006-02-13 | 68.5 | 26.8292 | IMB | hvtrip, fluctuating uAmp
37 | 13 | W4P04P34A | 21118 | 4.1.15.07 | 4.1.02 | 1720 |  | 2004-05-27 | 2005-12-19 | SA4140 | 1764 | -44 | 2006-02-13 | 61.6 | 26.7221 | IMB | hvtrip
38 | 13 | W4P04P36A | 21186 | 4.2.01.03 | 4.1.08 | 1640 |  | 2004-11-18 | 2005-12-19 | SA4107 | 1770 | -130 | 2006-02-13 | 58.5 | 26.8342 | IMB | flasher candidate / hvtrip
39 | 12 | W2P03P10A | 20318 | 2.1.01.11 | 2.0.00 | 2080 | 200 | 2005-12-26 | 2005-12-27 | SA4212 | 1883 | 197 | 2006-02-13 | 73.9 | 26.9461 | IMB | hvtrip (new!)
40 | 12 | W2P03P13B | 20437 | 2.1.10.06 | 2.0.09 | 1940 | 200 | 2006-02-17 |  | SA4213 | 1945 | -5 |  |  |  | Rec | fluctuating current ~300 uAmp
41 | 12 | W2P03P16B | 20539 | 2.1.18.12 | 2.1.05 | 2120 |  | 2002-11-26 | 2005-12-27 | SA4157 | 2123 | -3 | 2006-02-13 | 75.5 | 26.8795 | IMB | hvtrip
42 | 12 | W2P03P18B | 20607 | 2.2.04.08 | 2.1.11 | 1815 | 400 | 2005-11-01 | 2005-12-27 | SA4162 | 1822 | -7 | 2006-02-13 | 64.6 | 26.896 | IMB | no signal (was low occupancy in 9/05 already)
43 | 12 | W3P03P21B | 20709 | 3.1.06.08 | 3.0.05 | 2030 |  | 2005-12-16 | 2005-12-27 | SA4238 | 2029 | 1 | 2006-02-13 | 72.7 | 26.723 | Upg | hvtrip (new!)
44 | 12 | W3P03P23A | 20760 | 3.1.12.03 | 3.0.11 | 1920 |  | 2003-02-20 | 2005-12-27 | SA4214 | 1916 | 4 | 2006-02-13 | 68.5 | 26.8292 | Upg | hvtrip
45 | 12 | W3P03P27A | 20896 | 3.2.03.07 | 3.1.10 | 1920 | 200 | 2005-12-16 | 2005-12-27 | SA4155 | 1912 | 8 | 2006-02-13 | 68.5 | 26.8292 | IMB | hvtrip (new!)
46 | 12 | W4P03P34A | 21117 | 4.1.15.05 | 4.1.02 | 1720 |  | 2006-02-17 |  | SA4146 | 1822 | -102 |  |  | - | Upg | fluctuating current ~500 uAmp
47 | 11 | W1P02P06B | 20198 | 1.1.16.12 | 1.1.03 | 1690 | 150 | 2004-10-04 | 2006-01-06 | SA4199 | 1793 | -103 | 2006-02-13 | 60.5 | 26.7339 | IMB | flasher; killed QTC daughtercard, too
48 | 11 | W2P02P10B | 20334 | 2.1.01.10 | 2.0.00 | 2080 |  | 2005-04-06 | 2006-01-06 | SA4160 | 2055 | 25 | 2006-02-13 | 73.9 | 26.9461 | Rec | no signal (after blasting), HV channel ok (880 uA)
49 | 11 | W2P02P17A | 20555 | 2.2.01.07 | 2.1.08 | 2010 | 200 | 2006-02-08 | 2006-02-09 | SA4068 | 2019 | -9 |  |  | - | IMB | hvtrip (new!)
50 | 11 | W4P02P32B | 21065 | 4.1.09.08 | 4.0.08 | 2040 |  | 2002-08-14 | 2006-01-06 | SA4254 | 2042 | -2 | 2006-02-13 | 73 | 26.7452 | - | Bad Cable - Capped
51 | 11 | W4P02P34B | 21133 | 4.1.15.04 | 4.1.02 | 1720 |  | 2004-01-07 | 2006-01-06 | SA4220 | 1817 | -97 | 2006-02-13 | 61 | 26.9967 | IMB | hvtrip
52 | 10 | W3P01P21A | 20690 | 3.1.06.03 | 3.0.05 | 2030 |  | 2003-11-25 | 2006-01-12 | SA3705 | 2036 | -6 | 2006-02-13 | 72.7 | 26.723 | Upg | hvtrip
53 | 10 | W3P01P24B | 20809 | 3.1.14.10 | 3.1.01 | 1950 |  | 2002-08-23 | 2006-01-12 | SA3711 | 1966 | -16 | 2006-02-13 | 70.2 | 26.5778 | - | Bad Cable - Capped
54 | 10 | W4P01P33A | 21081 | 4.1.12.03 | 4.0.11 | 1860 |  | 2002-11-22 | 2006-01-12 | SA4233 | 1871 | -11 | 2006-02-13 | 66.2 | 26.8967 | Rec | hvtrip
55 | 10 | W4P01P34B | 21132 | 4.1.15.02 | 4.1.02 | 1720 |  | 2003-01-01 | 2006-01-12 | SA4192 | 1776 | -56 | 2006-02-13 | 61 | 26.9967 | Rec | hvtrip
56 | 10 | W4P01P38B | 21268 | 4.2.06.06 | 4.2.01 | 1940 |  | 2005-10-17 | 2006-01-12 | SA3911 | 1953 | -13 | 2006-02-13 | 69.4 | 26.7539 | Rec | no signal (was low occupancy in 9/05 already)
57 | 8 | W3M01P26B | 20875 | 3.1.20.02 | 3.1.07 | 1960 |  | 2002-12-06 | 2006-01-21 | SA3975 | 1969 | -9 |  |  | - | Rec | hvtrip
58 | 8 | W4M01P34A | 21113 | 4.1.14.09 | 4.1.01 | 1815 |  | 2006-02-17 |  | SA4112 | 1816 | -1 |  |  | - | Rec | fluctuating current ~150 uAmp
59 | 7 | W2M02P11A | 20347 | 2.1.03.11 | 2.0.02 | 2030 |  | 2006-02-17 |  | SA4178 | 2029 | 1 |  |  | - | Rec | fluctuating current ~150 uAmp
60 | 7 | W2M02P15B | 20500 | 2.1.15.04 | 2.1.02 | 1905 |  | 2003-03-15 | 2006-01-26 | SA4136 | 1905 | 0 |  |  | - | Upg | hvtrip
61 | 7 | W2M02P17B | 20568 | 2.1.20.12 | 2.1.07 | 1830 |  | 2005-06-15 | 2006-01-26 | SA4150 | 1831 | -1 |  |  | - | Rec | no signal (slow death?), HV channel ok (780 uA)
62 | 7 | W3M02P22B | 20738 | 3.1.08.08 | 3.0.07 | 2115 | 51 | 2004-11-13 | 2006-01-26 | SA3228 | 2095 | 20 |  |  | - | Rec | hvtrip
63 | 7 | W4M02P30B | 20993 | 4.1.03.04 | 4.0.02 | 1875 |  | 2005-04-07 | 2006-01-26 | SA4188 | 1886 | -11 |  |  | - | Rec | no signal (after blasting), HV channel ok (807 uA)
64 | 7 | W4M02P34A | 21112 | 4.1.14.07 | 4.1.01 | 1815 |  | 2005-04-14 | 2006-01-26 | SA4217 | 1823 | -8 |  |  | - | Rec | hvtrip
65 | 6 | W1M03P09A | 20278 | 1.2.04.07 | 1.1.11 | 2040 | 51 | 2004-09-29 | 2006-01-31 | SA3542 | 2041 | -1 |  |  | - | Rec | flasher candidate
66 | 6 | W4M03P36B | 21196 | 4.1.20.02 | 4.1.07 | 1940 |  | 2002-12-14 | 2006-01-31 | SA4184 | 1945 | -5 |  |  | - | Rec | flasher candidate
67 | 5 | W2M04P12A | 20379 | 2.1.06.05 | 2.0.05 | 1995 |  | 2002-11-12 | 2006-02-04 | SA3976 | 2008 | -13 |  |  | - | Upg | hvtrip, fluctuating uAmp
68 | 4 | W1M05P08B | 20259 | 1.2.01.06 | 1.1.08 | 2000 |  | 2002-11-11 | 2006-02-09 | SA3946 | 2014 | -14 |  |  | - | Upg | hvtrip, short > 1500V
69 | 4 | W2M05P12B | 20395 | 2.1.06.04 | 2.0.05 | 1995 | 150 | 2003-06-16 | 2006-02-09 | SA3714 | 1962 | 33 |  |  | - | Rec | comes and goes, very faint
70 | 3 | W1M06P05B | 20156 | 1.1.12.10 | 1.0.11 | 1760 |  | 2005-12-08 | 2006-02-14 | SA4209 | 1777 | -17 |  |  | - | Rec | hvtrip
71 | 3 | W3M06P26A | 20853 | 3.1.19.03 | 3.1.06 | 1925 | 100 | 2004-12-20 | 2006-02-14 | SA4145 | 1907 | 18 |  |  | - | Rec | no signal (base resistance ok, though)
72 | 3 | W4M06P38A | 21244 | 4.2.05.03 | 4.2.00 | 1930 | 51 | 2004-04-26 | 2006-02-14 | SA4191 | 1935 | -5 |  |  | - | Rec | mystery tube: no signal, slowly died…
73 | 2 | W3M07P24B | 20801 | 3.1.13.06 | 3.1.00 | 1940 |  | 2002-12-15 | 2006-02-18 | SA4152 | 1949 | -9 |  |  | - | Upg | hvtrip
74 | 2 | W3M07P25B | 20835 | 3.1.16.04 | 3.1.03 | 1905 |  | 2002-11-17 | 2006-02-18 | SA4194 | 1899 | 6 |  |  | - | Upg | hvtrip
75 | 2 | W4M07P31A | 21005 | 4.1.05.03 | 4.0.04 | 1900 | 51 | 2004-04-23 | 2006-02-18 | SA4169 | 1897 | 3 |  |  | - | Rec | mystery tube: no signal, slowly died…
76 | 2 | W4M07P37B | 21226 | 4.2.02.04 | 4.1.09 | 1890 |  | 2004-10-12 | 2006-02-18 | SA4114 | 1894 | -4 |  |  | - | Rec | came and went since 4/04, now signal too faint
77 | 2 | W4M07P38B | 21260 | 4.2.05.02 | 4.2.00 | 1930 |  | 2004-07-10 | 2006-02-18 | SA4234 | 1900 | 30 |  |  | - | Upg | hvtrip
78 | 1 | W1M08P08B | 20256 | 1.1.20.12 | 1.1.07 | 1630 |  | 2002-12-28 |  | SA4245 | 1781 | -151 |  |  | - | Rec | no signal, base resistance <20M
79 | 1 | W3M08P20B | 20664 | 3.1.01.12 | 3.0.00 | 1615 | 100 | 2003-02-11 |  | SA4243 | 1783 | -168 |  |  | - | Rec | flasher
80 | 1 | W3M08P23A | 20749 | 3.1.10.05 | 3.0.09 | 2060 | 350 | 2002-12-11 |  | SA4252 | 2052 | 8 |  |  | - | Rec | hvtrip (arcing relay repaired)
81 | 1 | W3M08P29A | 20953 | 3.2.07.05 | 3.2.02 | 1710 | 100 | 2003-12-20 |  | SA4197 | 1807 | -97 |  |  | - | Rec | hvtrip
82 | 0 | B1P04P02B | 21791 | 1.2.16.02 | 1.2.11 | 1790 |  | 2004-09-09 |  | SA4246 | 1799 | -9 |  |  | - | Rec | hvtrip
83 | 0 | B1P06P04A | 21838 | 1.2.18.12 | 1.3.01 | 1815 |  | 2003-01-15 |  | SA4244 | 1811 | 4 |  |  | - | Rec | hvtrip
84 | 0 | B2M01P04B | 21845 | 2.2.18.01 | 2.3.01 | 1890 |  | 2005-10-03 |  | SA4210 | 1839 | 51 |  |  | - | Rec | hvtrip
85 | 0 | B2M01P05B | 21868 | 2.2.19.02 | 2.3.02 | 1895 |  | 2004-07-18 |  | SA4216 | 1897 | -2 |  |  | - | Rec | hvtrip
86 | 0 | B2M02P01B | 21754 | 2.2.14.04 | 2.2.09 | 1870 |  | 2002-09-28 |  | SA4218 | 1871 | -1 |  |  | - | - | Bad Cable - Capped
87 | 0 | B2M02P02B | 21786 | 2.2.15.08 | 2.2.10 | 1870 |  | 2005-07-22 |  | SA4151 | 1875 | -5 |  |  | - | Rec | hvtrip
88 | 0 | B2M03P01B | 21753 | 2.2.14.06 | 2.2.09 | 1870 |  | 2005-12-26 |  | SA4153 | 1872 | -2 |  |  | - | Rec | hvtrip
89 | 0 | B2M05P04A | 21828 | 2.2.18.08 | 2.3.01 | 1890 |  | 2004-10-03 |  | SA4250 | 1841 | 49 |  |  | - | Rec | hvtrip
90 | 0 | B2M06P02B | 21782 | 2.2.16.04 | 2.2.11 | 1880 |  | 2003-01-04 |  | SA4219 | 1876 | 4 |  |  | - | Rec | hvtrip
91 | 0 | B2M06P05A | 21852 | 2.2.19.11 | 2.3.02 | 1895 |  | 2006-02-17 |  | SA4158 | 1895 | 0 |  |  | - | Rec | fluctuating current ~150 uAmp
92 | 0 | B3M01M03A | 21644 | 3.2.16.09 | 3.2.11 | 1960 |  | 2004-11-03 |  | SA4167 | 1951 | 9 |  |  | - | Rec | hvtrip
93 | 0 | B3M01M04B | 21630 | 3.2.18.01 | 3.3.01 | 1960 |  | 2003-02-22 |  | SA4182 | 1952 | 8 |  |  | - | Rec | no signal, base resistance <20M
94 | 0 | B3M02M06B | 21584 | 3.2.20.03 | 3.3.03 | 2000 |  | 2002-09-16 |  | SA4253 | 2011 | -11 |  |  | - | - | Bad Cable - Capped
95 | 0 | B3M04M04A | 21614 | 3.2.18.06 | 3.3.01 | 1960 |  | 2005-10-28 |  | SA4164 | 1953 | 7 |  |  | - | Rec | hvtrip, then 600 uAmp
96 | 0 | B4P01M01A | 21708 | 4.2.14.02 | 4.2.09 | 1825 |  | 2004-09-12 |  | SA4171 | 1845 | -20 |  |  | - | Rec | hvtrip
97 | 0 | B4P01M03B | 21660 | 4.2.16.11 | 4.2.11 | 1995 |  | 2002-09-28 |  | SA4215 | 1998 | -3 |  |  | - | - | Bad Cable - Capped
98 | 0 | B4P02M01A | 21709 | 4.2.14.04 | 4.2.09 | 1825 |  | 2003-02-13 |  | SA4147 | 1847 | -22 |  |  | - | Rec | hvtrip
99 | 0 | B4P02M02A | 21677 | 4.2.15.08 | 4.2.10 | 1965 |  | 2002-09-28 |  | SA4248 | 1965 | 0 |  |  | - | - | Bad Cable - Capped
100 | 0 | B4P02M06B | 21587 | 4.2.20.04 | 4.3.03 | 2105 |  | 2004-02-28 |  | SA4165 | 2111 | -6 |  |  | - | Rec | hvtrip
101 | 0 | B4P04M01A | 21711 | 4.2.14.08 | 4.2.09 | 1825 |  | 2003-01-02 |  | SA4202 | 1847 | -22 |  |  | - | Rec | no signal, base resistance <20M
102 | 0 | B4P07M04B | 21637 | 4.2.19.01 | 4.3.02 | 2090 |  | 2002-09-28 |  | SA4249 | 2039 | 51 |  |  | - | - | Bad Cable - Capped
color codes:
|  | installed 1995/6, IMB tube | [5 empty] | found bad or dead during 2005/6 reconstruction dark tests
|  | installed 2001, upgrade |  |  |  |  | remove zener |  | not installed yet |  |  | add zener
|  | installed 2002, reconstruction phase I | [7 empty] | 56m cable |  | increase HV, then add zeners to the remaining channels on same paddle card
bad cable, capped in 2002

## operating voltages
Date for selecting operating voltages for the Upgrade 2005 replacement PMTs
(Kate and Katie)
Set
[5 empty] | Voltage | [5 empty] | Voltages in 50 volt increments starting at 1400 V
| Tube | Date |  | HV for | HV for | Installed
|  |  |  | 30 mv | 25 mv | Address
New 2005 Upgrade PMTs
1 | SA4102 | 38561 | KW | 1660 | 1613 | 2.2.08.01 | 0 | 0 | 13.5 | 18 | 22.5 | 28.5 | 35 | 40.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2 | SA4103 | 38562 | KW | 1800 | 1747 | 2.1.16.08 | 0 | 0 | 0 | 0 | 0 | 15.5 | 20 | 24.5 | 30 | 35 | 43.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
3 | SA4104 | 38558 | KW | 1820 | 1746 | 2.1.16.12 | 0 | 0 | 0 | 0 | 0 | 17.5 | 21.5 | 25 | 28.5 | 32.5 | 40.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
4 | SA4107 | 38558 | KW | 1825 | 1769 | 4.2.01.03 |  | 0 | 0 | 0 | 0 | 0 | 18.5 | 20 | 28.5 | 35 | 45 | 47.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | SA4108 | 38558 | KW | 1775 | 1695 | 4.2.10.10 | 0 | 0 | 0 | 0 | 18 | 20.5 | 25 | 28 | 36 | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
6 | SA4109 | 38556 | KS | 2110 | 2067 | 3.1.10.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 20.5 | 23 | 29.5 | 31.5 | 35 | 0 | 0
7 | SA4110 | 38559 | KW | 1900 | 1843 | 3.2.09.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 19.5 | 25 | 30 | 37.5 | 43 | 0 | 0 | 0 | 0 | 0 | 0
8 | SA4111 | 38556 | KS | 1870 | 1820 | 1.2.08.09 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 18 | 22 | 29 | 32 | 37.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
9 | SA4112 | 38559 | KW | 1865 | 1816 | 4.1.14.09 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 19 | 22.5 | 27 | 32.5 | 39.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
10 | SA4114 | 38559 | KW | 1980 | 1894 | 4.2.02.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 22 | 27.5 | 28.5 | 33.5 | 41 | 0 | 0 | 0 | 0 | 0
11 | SA4115 | 38561 | KW | 1915 | 1860 | 4.1.12.12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 19.5 | 23 | 28.5 | 33.5 | 38 | 0 | 0 | 0 | 0 | 0 | 0
12 | SA4116 | 38559 | KW | 2075 | 2014 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 19 | 23.5 | 28 | 32.5 | 37 | 0 | 0 | 0
13 | SA4119 | 38558 | KW | 1800 | 1754 | 2.1.13.06 |  | 0 | 0 | 0 | 0 | 15 | 19.5 | 23 | 30 | 34 | 42.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
14 | SA4120 | 38558 | KW | 1800 | 1727 | 4.2.01.08 | 0 | 0 | 0 | 0 | 0 | 19 | 21.5 | 29 | 30 | 36.5 | 48.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
15 | SA4121 | 38559 | KW | 1870 | 1800 | 1.1.03.03 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 20 | 24.5 | 28.5 | 34.5 | 39.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
16 | SA4126 | 38561 | KW | 1760 | 1707 | 3.1.01.08 | 0 | 0 | 0 | 0 | 14.5 | 18.5 | 23 | 29 | 35 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
17 | SA4127 | 2005-07-25 | KW | 1640 | 1576 | 3.2.11.03 |  | 14 | 17.5 | 21.5 | 27.5 | 31.5 | 37.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
18 | SA4131 | 38556 | KS | 1675 | 1649 | 3.2.12.05 | 0 | 0 | 0 | 15 | 18.5 | 24 | 31.5 | 35 | 44.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
19 | SA4132 | 38559 | KW | 1800 | 1757 | 3.2.08.12 |  | 0 | 0 | 0 | 0 | 14.5 | 18.5 | 23.5 | 30 | 34.5 | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
20 | SA4133 | 38558 | KW | 1780 | 1711 | 3.2.07.03 | 0 | 0 | 0 | 0 | 0 | 20.5 | 23 | 26.5 | 36 | 42.5 | 48.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
21 | SA4136 | 38561 | KW | 1960 | 1905 | 2.1.15.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 19 | 25 | 29 | 33.5 | 38.5 | 0 | 0 | 0 | 0 | 0
22 | SA4137 | 38558 | KW | 1780 | 1707 | 1.2.03.05 | 0 | 0 | 0 | 0 | 15.5 | 19 | 23 | 26.5 | 37 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
23 | SA4140 | 38559 | KW | 1830 | 1764 | 4.1.15.07 |  | 0 | 0 | 0 | 0 | 16.5 | 17.5 | 22.5 | 26.5 | 35 | 38.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
24 | SA4141 | 38559 | KW | 1765 | 1709 | 3.2.07.09 | 0 | 0 | 0 | 0 | 16 | 18 | 24 | 27.5 | 33.5 | 40.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
25 | SA4142 | 38556 | KS | 1750 | 1689 | 2.2.08.11 | 0 | 0 | 0 | 0 | 14.5 | 19.5 | 26 | 30 | 40.5 | 53 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
26 | SA4143 | 38561 | KW | 1920 | 1867 | 1.1.05.12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 18 | 22.5 | 28 | 34 | 38 | 0 | 0 | 0 | 0 | 0 | 0
27 | SA4144 | 38559 | KW | 2020 | 1972 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 19 | 23.5 | 28 | 32.5 | 37 | 0 | 0 | 0
28 | SA4145 | 38559 | KW | 1965 | 1907 | 3.1.19.03 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 21 | 24 | 28 | 32.5 | 37.5 | 0 | 0 | 0 | 0 | 0
29 | SA4146 | 38558 | KW | 1900 | 1822 | 4.1.15.05 |  | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 23.5 | 27 | 30 | 35 | 41 | 0 | 0 | 0 | 0 | 0 | 0
30 | SA4147 | 38556 | KS | 1900 | 1845 | 4.2.14.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 21 | 25 | 30 | 35 | 45 | 0 | 0 | 0 | 0 | 0 | 0
31 | SA4147 | 38561 | KW | 1900 | 1847 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 20 | 24.5 | 30 | 35 | 42.5 | 0 | 0 | 0 | 0 | 0 | 0 | duplicate entry!!
32 | SA4149 | 38558 | KW | 1830 | 1775 | 1.1.14.07 | 0 | 0 | 0 | 0 | 0 | 13.5 | 17 | 23 | 27.5 | 32 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
33 | SA4150 | 38558 | KW | 1880 | 1831 | 2.1.20.12 | 0 | 0 | 0 | 0 | 0 | 0 | 13.5 | 18 | 21.5 | 26 | 32 | 35.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
34 | SA4151 | 38558 | KW | 1950 | 1875 | 2.2.15.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 22 | 22.5 | 25 | 30 | 34 | 40.5 | 0 | 0 | 0 | 0 | 0
35 | SA4152 | 38556 | KS | 2000 | 1949 | 3.1.13.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 19.5 | 23 | 30 | 36 | 43 | 0 | 0 | 0 | 0
36 | SA4153 | 38558 | KW | 1950 | 1872 | 2.2.14.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 23 | 26.5 | 30 | 38.5 | 43.5 | 0 | 0 | 0 | 0 | 0
37 | SA4154 | 38561 | KW | 1840 | 1792 | 3.1.18.10 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 20 | 25 | 31 | 36 | 42.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
38 | SA4155 | 38561 | KW | 1965 | 1912 | 3.2.03.07 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 19 | 23.5 | 28.5 | 33.5 | 37 | 0 | 0 | 0 | 0 | 0
39 | SA4157 | 38559 | KW | 2185 | 2123 | 2.1.18.12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 19 | 22 | 26 | 32 | 35 | 0
40 | SA4158 | 38558 | KW | 1970 | 1895 | 2.2.19.11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 22 | 25 | 28 | 33.5 | 40 | 0 | 0 | 0 | 0 | 0
41 | SA4159 | 38559 | KW | 1940 | 1882 | 4.1.04.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 22 | 26 | 31 | 36.5 | 43 | 0 | 0 | 0 | 0 | 0
42 | SA4160 | 38559 | KW | 2120 | 2055 | 2.1.01.10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 19 | 24.5 | 28.5 | 34.5 | 38 | 0 | 0
43 | SA4161 | 38561 | KW | 2050 | 1998 | 1.1.08.12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 20 | 26.5 | 30 | 35 | 38.5 | 0 | 0 | 0
44 | SA4162 | 38561 | KW | 1875 | 1822 | 2.2.04.08 | 0 | 0 | 0 | 0 | 0 | 0 | 14 | 17 | 22 | 27.5 | 33 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0
45 | SA4163 | 38556 | KS | 1850 | 1797 | 2.2.05.06 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 20 | 26.5 | 30 | 35 | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0
46 | SA4164 | 38561 | KW | 2010 | 1953 | 3.2.18.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 20 | 24.5 | 29 | 33.5 | 39 | 0 | 0 | 0 | 0
47 | SA4165 | 38559 | KW | 2170 | 2111 | 4.2.20.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 19 | 23.5 | 27 | 34 | 36 | 0
48 | SA4167 | 38558 | KW | 2015 | 1951 | 3.2.16.09 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 19 | 24 | 28.5 | 35.5 | 42 | 0 | 0 | 0 | 0
49 | SA4168 | 38561 | KW | 2065 | 2017 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 19.5 | 22.5 | 26.5 | 33 | 37 | 0 | 0 | 0
50 | SA4169 | 38558 | KW | 1980 | 1897 | 4.1.05.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 20.5 | 24.5 | 28.5 | 34.5 | 38.5 | 0 | 0 | 0 | 0 | 0
51 | SA4171 | 38561 | KW | 1890 | 1845 | 4.2.14.02 | 0 | 0 | 0 | 0 | 0 | 0 | 14 | 16 | 19.5 | 25.5 | 30.5 | 34 | 0 | 0 | 0 | 0 | 0 | 0 | 0
52 | SA4178 | 38558 | KW | 2095 | 2029 | 2.1.03.11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 22.5 | 28 | 31 | 34.5 | 39.5 | 0 | 0
53 | SA4179 | 38558 | KW | 2040 | 1984 | 3.2.01.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19 | 21.5 | 25 | 31 | 35.5 | 40 | 0 | 0 | 0
54 | SA4180 | 38561 | KW | 1990 | 1931 | 3.1.12.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 17.5 | 21 | 27 | 31 | 36 | 0 | 0 | 0 | 0 | 0
55 | SA4182 | 38558 | KW | 2020 | 1952 | 3.2.18.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 19.5 | 23.5 | 28.5 | 35 | 40 | 0 | 0 | 0 | 0
56 | SA4184 | 38562 | KW | 2000 | 1946 | 4.1.20.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 13.5 | 16 | 21 | 25 | 30 | 34 | 0 | 0 | 0 | 0 | 0
57 | SA4188 | 38559 | KW | 1950 | 1886 | 4.1.03.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 20.5 | 27 | 30 | 36 | 39.5 | 0 | 0 | 0 | 0 | 0
58 | SA4191 | 38558 | KW | 2000 | 1935 | 4.2.05.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 22 | 26.5 | 30 | 36.5 | 40.5 | 0 | 0 | 0 | 0
59 | SA4192 | 38558 | KW | 1850 | 1776 | 4.1.15.02 | 0 | 0 | 0 | 0 | 0 | 0 | 19.5 | 23 | 26 | 30 | 37 | 42 | 0 | 0 | 0 | 0 | 0 | 0 | 0
60 | SA4194 | 38559 | KW | 1960 | 1899 | 3.1.16.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 20 | 24.5 | 29 | 35.5 | 41 | 0 | 0 | 0 | 0 | 0
61 | SA4195 | 38558 | KW | 2080 | 1998 | 1.1.08.09 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19.5 | 21.5 | 24 | 28.5 | 32 | 35 | 0 | 0 | 0
62 | SA4197 | 38561 | KW | 1865 | 1807 | 3.2.07.05 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 18 | 23 | 28 | 35.5 | 42.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
63 | SA4198 | 38558 | KW | 2075 | 2023 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 19 | 22.5 | 27.5 | 32 | 34 | 0 | 0 | 0
64 | SA4199 | 38558 | KW | 1875 | 1793 | 1.1.16.12 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 20.5 | 25.5 | 28 | 36.5 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0
65 | SA4202 | 38561 | KW | 1905 | 1847 | 4.2.14.08 | 0 | 0 | 0 | 0 | 0 | 0 | 13 | 15.5 | 19 | 25.5 | 29 | 35.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
66 | SA4203 | 38558 | KW | 1740 | 1699 | 4.2.10.04 | 0 | 0 | 0 | 0 | 16 | 18.5 | 24 | 31 | 36 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
67 | SA4205 | 38559 | KW | 1740 | 1687 | 2.2.10.04 |  | 0 | 0 | 0 | 0 | 21.5 | 26 | 32 | 37 | 45.5 | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
68 | SA4206 | 38559 | KW | 1750 | 1700 | 3.2.10.10 | 0 | 0 | 0 | 0 | 14 | 19.5 | 25 | 30 | 37 | 41.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
69 | SA4207 | 38558 | KW | 1710 | 1648 | 4.2.08.01 |  | 0 | 0 | 16.5 | 19 | 24.5 | 29 | 35 | 44.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
70 | SA4209 | 38558 | KW | 1830 | 1777 | 1.1.12.10 | 0 | 0 | 0 | 0 | 0 | 14 | 17.5 | 22 | 26.5 | 32 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
71 | SA4210 | 38559 | KW | 1900 | 1839 | 2.2.18.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 20.5 | 26.5 | 30 | 35.5 | 39.5 | 0 | 0 | 0 | 0 | 0 | 0
72 | SA4211 | 38558 | KW | 1870 | 1812 | 1.1.17.10 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 20 | 22.5 | 28 | 33 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0
73 | SA4212 | 38556 | KS | 1965 | 1883 | 2.1.01.11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19.5 | 22 | 25.5 | 29 | 35 | 44.5 | 0 | 0 | 0 | 0 | 0
74 | SA4213 | 38561 | KW | 2000 | 1945 | 2.1.10.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 20.5 | 26 | 30 | 33.5 | 40 | 0 | 0 | 0 | 0
75 | SA4214 | 38559 | KW | 2000 | 1916 | 3.1.12.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 21.5 | 24 | 27.5 | 30 | 36.5 | 46.5 | 0 | 0 | 0 | 0
76 | SA4215 | 38558 | KW | 2060 | 1998 | 4.2.16.11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 21 | 25 | 29 | 33.5 | 39.5 | 0 | 0 | 0 | <== was mixed up with SA4251
77 | SA4216 | 38559 | KW | 1965 | 1897 | 2.2.19.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 22 | 24 | 28 | 35 | 39.5 | 0 | 0 | 0 | 0 | 0
78 | SA4217 | 38556 | KS | 1880 | 1823 | 4.1.14.07 | 0 | 0 | 0 | 0 | 0 | 0 | 14 | 18.5 | 21 | 26 | 32 | 39.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
79 | SA4218 | 38558 | KW | 1940 | 1871 | 2.2.14.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 19 | 22.5 | 27.5 | 32 | 35 | 0 | 0 | 0 | 0 | 0 | 0
80 | SA4219 | 38558 | KW | 1925 | 1876 | 2.2.16.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 18.5 | 21.5 | 26 | 31.5 | 36 | 0 | 0 | 0 | 0 | 0 | 0
81 | SA4220 | 38558 | KW | 1880 | 1818 | 4.1.15.04 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 19 | 23 | 27 | 31 | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0
82 | SA4233 | 38561 | KW | 1915 | 1871 | 4.1.12.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 17 | 20.5 | 26.5 | 33.5 | 40.5 | 0 | 0 | 0 | 0 | 0 | 0
83 | SA4234 | 38559 | KW | 1960 | 1900 | 4.2.05.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 20 | 24 | 29 | 33 | 42 | 0 | 0 | 0 | 0 | 0
84 | SA4235 | 38559 | KW | 2070 | 2021 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 17.5 | 23.5 | 28 | 31.5 | 35.5 | 0 | 0 | 0
85 | SA4237 | 38561 |  | 1970 | 1919 | 3.1.12.08 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 17.5 | 22.5 | 28 | 33.5 | 37.5 | 0 | 0 | 0 | 0 | 0
86 | SA4238 | 38559 | KW | 2080 | 2029 | 3.1.06.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 17.5 | 20.5 | 26.5 | 32 | 36.5 | 0 | 0 | 0
87 | SA4241 | 38558 | KW | 1925 | 1862 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 19.5 | 22.5 | 27 | 33.5 | 38.5 | 0 | 0 | 0 | 0 | 0 | 0
88 | SA4243 | 38558 | KW | 1850 | 1783 | 3.1.01.12 |  | 0 | 0 | 0 | 0 | 0 | 18 | 22 | 26.5 | 30 | 35.5 | 41.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
89 | SA4244 | 38558 | KW | 1865 | 1811 | 1.2.18.12 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 18 | 22 | 28 | 34 | 40 | 0 | 0 | 0 | 0 | 0 | 0 | 0
90 | SA4245 | 38556 | KS | 1850 | 1781 | 1.1.20.12 |  | 0 | 0 | 0 | 0 | 0 | 17.5 | 22 | 26 | 30 | 39 | 48 | 0 | 0 | 0 | 0 | 0 | 0 | 0
91 | SA4246 | 38561 | KW | 1850 | 1799 | 1.2.16.02 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 19.5 | 25 | 30 | 34 | 38.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0
92 | SA4248 | 38559 | KW | 2025 | 1965 | 4.2.15.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 20.5 | 23 | 26.5 | 31 | 36.5 | 0 | 0 | 0 | 0
93 | SA4249 | 38559 | KW | 2130 | 2039 | 4.2.19.01 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19.5 | 22.5 | 25 | 28.5 | 32 | 38.5 | 0 | 0
94 | SA4250 | 38558 | KW | 1900 | 1841 | 2.2.18.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 21 | 24.5 | 30 | 35 | 40 | 0 | 0 | 0 | 0 | 0 | 0
95 | SA4251 | 38559 | KW | 2035 | 1985 | 4.1.18.12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 17 | 20.5 | 26 | 32 | 34 | 0 | 0 | 0 | 0 | <== actually installed, not SA4215
96 | SA4252 | 38559 | KW | 2115 | 2052 | 3.1.10.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 20.5 | 24 | 28 | 34.5 | 39.5 | 0 | 0
97 | SA4253 | 38559 | KW | 2065 | 2011 | 3.2.20.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16 | 19.5 | 25 | 28.5 | 32.5 | 35.5 | 0 | 0 | 0
98 | SA4254 | 38559 | KW | 2100 | 2042 | 4.1.09.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 22.5 | 25.5 | 30 | 32 | 36 | 0 | 0
99 | SA4255 | 38559 | KW | 2080 | 2030 | 3.1.06.09 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 21 | 26.5 | 32.5 | 37 | 42 | 0 | 0
100 | SA4256 | 38558 | KW | 2010 | 1948 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 21.5 | 26 | 29 | 32.5 | 37.5 | 0 | 0 | 0 | 0
Older PMTs (most should be usable in SK)
(i.e., new generation Hamamatsu tubes from a previous upgrade/reconstruction. Old tube ≠ IMB tube)
Set
Voltage
| Tube | Date |  | HV for | HV for | [20 empty] | Notes
|  |  |  | 30 mv | 25 mv
- | 9003 | 38560 | KW |  | 247 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | No signal
- | SA4004 | 38560 | KW |  | 247 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | No signal. Tube seal may be bad
1 | 9009 | 38561 | KW | 1600 | 1552 |  | 0 | 14.5 | 18.5 | 25 | 30 | 34.5 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, cable ~1.5m, sharp spe peak Q4 P4M5B or W4 P00P30A, 1585V on box
2 | SA2869 | 38560 | KW | 1615 | 1570 |  | 0 | 13 | 16 | 21.5 | 27.5 | 35 | 40.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 37 m cable, old tube, 1518V on box
3 | SA3288 | 38560 | KW | 1660 | 1602 |  | 0 | 0 | 15 | 19 | 24.5 | 28.5 | 35.5 | 42.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Dropped from truck, old tube, 37 m cable, 1589V on box
4 | 9010 | 38561 | KW | 1665 | 1604 |  | 0 | 0 | 15.5 | 19 | 25 | 28.5 | 33 | 42.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, 1500V on box, cable ~1.5 m
5 | SA2972 | 38560 | KW | 1700 | 1636 |  | 0 | 0 | 0 | 16 | 21 | 27 | 30 | 37 | 53.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 37 m cable, old tube, 1518V on box
6 | SA2002 | 38561 | KW | 1700 | 1641 |  | 0 | 0 | 0 | 16 | 21.5 | 26.5 | 30 | 34 | 43.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, SHV, Q2M2P5B or W4P3P36B, 1635V on box, cable ~ 2 ft
7 | SA3581 | 38559 | KW | 1710 | 1641 |  | 0 | 0 | 0 | 15.5 | 21 | 25.5 | 29 | 38.5 | 44 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 37 m cable, old tube; signal flickering but no apparent light leak; 1529V on box
8 | SA2000 | 38560 | KW | 1700 | 1650 |  | 0 | 0 | 0 | 14.5 | 18.5 | 24 | 30 | 36.5 | 45.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | short cable (1m?), narrow pe signal, old pmt black sheet on pmt, 1600V on box, also labeled T2M04B03B
9 | 9004 | 38561 | KW | 1715 | 1664 |  | 0 | 0 | 0 | 15.5 | 18 | 23.5 | 28 | 33 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, 1500 V on box, W?M6?34B removed from wall to be tested
10 | SA1622 | 38561 | KW | 1725 | 1671 |  | 0 | 0 | 0 | 13.5 | 19 | 22.5 | 28 | 32 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, SHV, Q3M6M2B or W2P2P31A, 1700V on box, cable~1 m
11 | SA2897 | 38559 | KW | 1740 | 1673 |  | 0 | 0 | 0 | 14 | 19 | 22.5 | 26.5 | 32 | 37 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 37 m cable, old tube, 1519V on box
12 | SA1918 | 38561 | KW | 1735 | 1675 |  | 0 | 0 | 0 | 0 | 17.5 | 21.5 | 27 | 33.5 | 40.5 | 44 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, W4P0P33B, 1655V on box, cable~1.5 m
13 | 9011 | 38623 | KW | 1740 | 1684 |  | 0 | 0 | 0 | 13 | 16.5 | 21.5 | 27 | 31 | 36 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, SHV, W1P2P5A or Q1P8P1B, 1695 V on box cable ~0.5 ft
14 | SA2053 | 38560 | KW | 1740 | 1689 |  | 0 | 0 | 0 | 13.5 | 16 | 18.5 | 24.5 | 32.5 | 36.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | P2P5B or Q3M5M5A, old tube, 1715V on box, SHV
15 | SA1919 | 38560 | KW | 1740 | 1691 |  | 0 | 0 | 0 | 12.5 | 16 | 18.5 | 25 | 30.5 | 37.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | W3P2P77B or Q3M7M2B, old tube, 1700V on box, SHV
16 | 0 | 38195 | KW | 1750 | 1702 |  | 0 | 0 | 0 | 0 | 0 | 19.5 | 25 | 30 | 35.5 | 43.5 | 50 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | MHV, in box with 2 pmts (short cable), 1654V on box; noise on scope, sharp spe peak
17 | 9008 | 38560 | KW | 1760 | 1711 |  | 0 | 0 | 0 | 0 | 14 | 17 | 23 | 28.5 | 34 | 43.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | WM4P35B, 1600V on box, SHV, 1ft cable
18 | SA1884 | 38561 | KW | 1765 | 1720 |  | 0 | 0 | 0 | 0 | 14 | 18 | 23 | 28 | 32 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube(?), SHV, W4P3P31A, 1705V on box
19 | 9006 | 38561 | KW | 1775 | 1722 |  | 0 | 0 | 0 | 0 | 14.5 | 17.5 | 21.5 | 27 | 32 | 39 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV WMS-28B, 1500 V on tube, cable ~6 in
20 | 9001 | 38561 | KW | 1800 | 1751 |  | 0 | 0 | 0 | 0 | 0 | 15 | 20 | 25 | 30 | 34 | 40.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, cable~6 in, No V on box
21 | SA1983 | 38561 | KW | 1820 | 1761 |  | 0 | 0 | 0 | 0 | 0 | 14 | 18.5 | 23.5 | 28.5 | 33.5 | 41 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, SHV, 1755V on box, cable ~1ft
22 | SA1966 | 38561 | KW | 1820 | 1771 |  | 0 | 0 | 0 | 0 | 0 | 15 | 17.5 | 21.5 | 28 | 32.5 | 37.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, 1750 V on box, W2M3P14A - removed
23 | 9007 | 38561 | KW | 1825 | 1778 |  | 0 | 0 | 0 | 0 | 0 | 14.5 | 17.5 | 21.5 | 27 | 31.5 | 36.5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, WM7P26B, 1500 V on box, cable ~ 1m
24 | SA1959 | 38561 | KW | 1865 | 1806 |  | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 20.5 | 23.5 | 28.5 | 32.5 | 38 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, W4P3P29B, 1820V on box, cable ~1m, much sharper pe peak
25 | 9005 | 38561 | KW | 1885 | 1823 |  | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 18 | 23.5 | 27.5 | 31 | 36 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | Old tube, MHV, sharper spe peak, WM5?1A?, 1700V on box, seal? cable ~1.5 m removed from wall to be tested
26 | SA3908 | 38560 | KW | 1935 | 1884 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 20.5 | 26.5 | 33 | 36 | 42.5 | 0 | 0 | 0 | 0 | 0 | old tube, 1742V on box, marked B3 M02M06B, removed, to be tested
27 | SA3930 | 38560 | KW | 1960 | 1888 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 22.5 | 25.5 | 29 | 34.5 | 39.5 | 0 | 0 | 0 | 0 | 0 | 56 m cable, old tube, 1875V on box
28 | SA3911 | 38560 | KW | 2020 | 1953 | 4.2.06.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 21 | 24.5 | 28 | 32.5 | 37.5 | 0 | 0 | 0 | 0 | 56 m cable, old tube, 1885V on box
29 | SA3714 | 38560 | KW | 2015 | 1962 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 20.5 | 24 | 27.5 | 32.5 | 37.5 | 0 | 0 | 0 | 0 | duplicate entry!!
30 | SA3714 | 38560 | KW | 2015 | 1962 | 2.1.06.04 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 20.5 | 24 | 27.5 | 32.5 | 37.5 | 0 | 0 | 0 | 0 | 56 m cable, old tube, 1885V on box
31 | SA3711 | 38560 | KW | 2040 | 1966 | 3.1.14.10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 19.5 | 22.5 | 27.5 | 32.5 | 37 | 0 | 0 | 0 | 0 | 56 m cable, old tube, 1866V on box
32 | SA3975 | 38560 | KW | 2030 | 1969 | 3.1.20.02 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 19 | 23 | 26.5 | 32.5 | 37.5 | 0 | 0 | 0 | 0 | 56 m cable, old tube, 1865V on box
33 | SA3976 | 38560 | KW | 2060 | 2008 | 2.1.06.05 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15.5 | 20 | 23 | 28.5 | 33 | 39.5 | 0 | 0 | 0 | 56 m cable, old tube, 1875V on box
34 | SA3946 | 38560 | KW | 2080 | 2014 | 1.2.01.06 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 18.5 | 22.5 | 26.5 | 32 | 40 | 0 | 0 | 0 | 56 m cable, old tube, 1875V on box
35 | SA4062 | 38559 | KW | 2070 | 2019 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 14.5 | 19.5 | 23 | 27 | 32 | 37.5 | 0 | 0 | 0 | 56 m cable, old tube, 1784V on box
36 | SA4068 | 38560 | KW | 2070 | 2019 | 2.2.01.07 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 19 | 23 | 27.5 | 31.5 | 37 | 0 | 0 | 0 | 56 m cable, old tube, 1875V on box
37 | SA4069 | 38560 | KW | 2090 | 2030 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18.5 | 22.5 | 26.5 | 30.5 | 35 | 39.5 | 0 | 0 | 56 m cable, old tube, 1854V on box, MHV very noisy signal, using pickoff box
38 | SA3705 | 38560 | KW | 2090 | 2036 | 3.1.06.03 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 21 | 26.5 | 31 | 37 | 41.5 | 0 | 0 | 56 m cable, old tube, 1895V on box
39 | SA3542 | 38560 | KW | 2105 | 2041 | 1.2.04.07 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 21 | 25 | 29 | 36 | 42 | 0 | 0 | 56 m cable, old tube, 1900V on box
40 | SA4067 | 38560 | KW | 2100 | 2042 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 20.5 | 25 | 30 | 35 | 39 | 0 | 0 | 56 m cable, old tube, 1774V on box, MHV, set with pickoff box
41 | SA4070 | 38560 | KW | 2110 | 2048 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 16.5 | 22 | 25.5 | 29 | 33 | 37 | 0 | 0 | 56 m cable, old tube, 1890V on box
42 | SA3796 | 38560 | KW | 2130 | 2055 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17.5 | 19.5 | 22.5 | 28 | 34.5 | 39 | 0 | 0 | 56 m cable, old tube, 1884V on box, MHV, set using pickoff box
43 | SA3709 | 38560 | KW | 2150 | 2072 |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 19.5 | 23.5 | 27 | 30 | 36.5 | 42.5 | 0 | 56 m cable, old tube
44 | SA3228 | 38560 | KW | 2155 | 2095 | 3.1.08.08 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 21 | 24.5 | 29 | 34.5 | 41 | 0 | 56 m cable, old tube, 1915V on box
[5 empty] | color code: | installed 1/31-2/18
missing in WRK's list
not installed yet (2/24-...)
duplicate entry